## Slide 1: DOG Cube Code
Documentation

## Slide 2: Overview

## Slide 3: Overview
This code is designed to control the entire DOG setup (Delivery On G-line, colloquially known as the Cube).
This involves interfacing with the Elveflow (gas flow and sheath flow control), pumps (sample/buffer flow control), valves (liquid path control), and SPEC (x-ray control and detector saving), as well as the local file system on the computer (config reading and logging)
Elveflow
Pumps + Valves
(Microcontroller)
SPEC
Local file system
GUI Code

## Slide 4: General usage
Start up the software
  Create the GUI
  Load the config (local file system)
  Start the logging (local file system)
  Connect to Elveflow and start up a thread to handle all Elveflow communication
  Connect to spec
Start microcontroller connection (pumps and valves)
Start sheath flow (Elveflow) and automatically graph
Elveflow
Pumps + Valves
SPEC
Local file system

## Slide 5: General usage
Press Auto Run
  Create a file name (spec)
  Double-check the Elveflow flowrates
  Start running everything, which involves switching valves, starting and stopping the pumps, telling spec to actually take scans, and graphing Elveflow results to the main page
  Clean everything, which again involves valves and pumps
Purge sheath if necessary (Elveflow and valves)
Elveflow
Pumps + Valves
SPEC
Local file system

## Slide 6: General usage
Shutdown
  Close connection to Elveflow (stop the sheath), spec, pump, and valve and the logging file (local file system)
Elveflow
Pumps + Valves
SPEC
Local file system

## Slide 7: Threads
There are a couple threads that we need to be careful of to manage them. Crashing any of them except the main thread will not throw an error and instead will just make the whole program, so this is very hard to debug!
  Main thread (the GUI runs here)
  Elveflow thread (queries the Elveflow every half second or so)
  Graph thread (updates the graph on the Elveflow tab periodically)
    Second graph on front page (?)
  Run queue (sends commands and spin-waits for responses from the pumps and the valves)
  Manual instrument thread (send manual commands to the pumps/valves)
  Spec thread (?)
  State diagram thread (?)

## Slide 8: Bugs
Initialize Sheath currently does not stop until it receives confirmation that it got to the right level—TODO: be allowed to cancel
Something going wrong with the Elveflow if you leave it running for too long

## Slide 9: Class: ElveflowHandler
hardware/FileIO.py

## Slide 10: ElveflowHandler
Two versions, primary version is ElveflowHandler_SDK (aliased as simply ElveflowHandler)
  The other one, ElveflowHandler_ESI, reads in data from a csv file coming from a separate instance of Elveflow ESI software.
  The _ESI version is very likely broken and not kept up.
A class that can send commands to and receive and store data from the Elveflow.

## Slide 11: Functions: initialization
__init__(self, sourcename, errorlogger, sensortypes)
  Initializes the connection to the Elveflow, and then initializes the connection to each of the sensors.
  Creates a data storage queue
    This is a queue.Queue, which should be thread-safe
  Writes warnings to the errorlogger if it cannot connect to the Elveflow or to the sensors

## Slide 12: Functions: read from machine
start(self)
  Starts a reading thread collecting data from the Elveflow sensor
    Every 0.1 seconds, read all four pressure and all four flow sensor readings into the data storage queue
    Writes a warning to the errorlogger if error in reading sensor (but just ignore errors from reading pressure)
  TODO: currently does not set run flag when it starts. Does not check if there already exists a reading thread
stop(self)
  Stops any currently reading thread or command thread
  If no currently active reading or command thread, silently does nothing

## Slide 13: Functions: store/manage data
fetchOne(self)
  Returns the oldest entry in the storage queue and removes it from the queue
  Returns None if the queue is empty
peekOne(self)
  Returns the oldest entry in the storage queue
  Returns None if the queue is empty
fetchAll(self)
  Returns all entries in the storage queue
getHeader(self)
  Returns the list of keys of the dicts that the fetch/peek functions return

## Slide 14: Functions: send commands
set_pressure_loop(self, channel_number, value, interrupt_event, on_finish)
  Starts a command thread sending instructions to set the pressure of channel channel_number to value
    Reads the current pressure (warn if cannot)
    In a loop: slowly increases/decreases in PRESSURE_MAXSLOPE increments every 0.1 s
    Writes a warning to the errorlogger if the pressure cannot be set
    When done, the thread automatically stops
    Runs on_finish() when the thread stops, whether naturally or because of an interrupt
  interrupt_event can be None, but it can also be a new Event which you can set to True to stop the thread from running from outside the class

## Slide 15: Functions: send commands
set_volume_loop(self, channel_number, value, interrupt_event, pid_constants)
  Starts a command thread sending instructions to set the flow rate of channel channel_number to value
    Run a PID loop with the PID constants (or default values if None)
    Writes a warning to the errorlogger if the pressure cannot be read or set
    Never stops until interrupted
    When interrupted, calls set_pressure_loop() to set the pressure to 0 (this creates a subthread!). Waits for this subthread to finish before exiting
  interrupt_event can be None, but it can also be a new Event which you can set to True to stop the thread from running from outside the class

## Slide 16: Functions: send commands (blocking)
run_volume(self, channel_number, value, interrupt_event, pid_constants, margin, stable_time, timeout)
  Like set_volume_loop, but does it in the calling thread (i.e. blocking).
  However, a few more parameters to allow for stopping on its own:
    Margin and stable_time: if we stay within margin flowrate for stable_time amount of time, stop. (margin is in absolute units, not relative percentages)
    Timeout: if we pass this amount of time, immediately stop

## Slide 17: Class: ElveflowDisplay
widgets/ElveflowDisplay.py

## Slide 18: ElveflowDisplay
Extends tk.Canvas
Represents and manages the entire Elveflow graphing tab
Interfaces between buttons/display elements and the ElveflowHandler instance

## Slide 20: Functions: initialization
__init__(self, window, height, width, elveflow_config, errorlogger, maingui, **kwargs)
  Parameters:
    window, height, width, and **kwargs are as in the parent tk.Canvas element
    elveflow_config is the part of the config file with information about the elveflow
    errorlogger and maingui are self-explanatory
  Set up internally used variables like the graph axis limits or locks
  Create the actual display: the buttons/inputs/empty graph

## Slide 21: Functions: updating
start(self)
  If an instance of ElveflowHandler is already running, raise error (not just write to errorlogger)
  Create an instance of ElveflowHandler and start the ElveflowHandler
  Start a new polling/display thread
    Every 0.5 s, grab data from the ElveflowHandler and call self.update_plot()
    Stop when the run_flag is set to False. Then, set the internal variable done_shutting_down to True or call stop()…?
    TODO: streamline this whole started_shutting_down and done_shutting_down flag business.
stop(self, shutdown)
  Stops the polling/display thread by setting the run_flag to False

## Slide 22: Functions: update
update_plot(self)
  Adjust plot x and y limits on the graph
  Actually update the graph data and call self.canvas.draw() to refresh the screen
    This is done in a spun up daemon thread. Apparently I had issues with shutdown when this wasn’t in a daemon thread? If time, investigate further.
  Update the sheath flow pressure display on the main tab. This is not where this should be done
populate_dropdowns(self)
  Populate the dropdowns for choosing x and y axes, based on the active ElveflowHandler

## Slide 23: Functions: data saving
start_saving(self)
  Create and open a new csv file
  Start saving data to it (set the save_flag to True; the actual saving happens in start())
  Writes error to errorlog if the ElveflowHandler header is not known yet
stop_saving(self, shutdown)
  Stop saving csv data (set the save_flag to False)
  Close the file

## Slide 24: Functions: button commands
set_axis_limits(self)
  Adjust plot x and y limits on the graph, then call update_plot()
start_pressure(self, channel, isPressure) and stop_pressure(self, channel)
  Start or stop the corresponding functions in the connected ElveflowHandler.
  Note that this is used for both pressure and volume control, depending on the isPressure parameter
  We hold onto the interrupt flag, which we can use to stop prematurely
run_volume(self, channel, target, margin, stable_time)
  Basically just a way to call the ElveflowHandler’s version of run_volume
  Remember, this is blocking!

## Slide 25: Class: Main
gui.py

## Slide 26: Main
The class itself does not extend tk.Tk (a tk window), but instead takes a preformed tk.Tk as an argument. Example usage:
  window = tk.Tk() Main(window) window.mainloop()
A class that creates the entire GUI when instantiated

## Slide 27: Functions: overview
This class is overstuffed by a whole lot. We really need to separate out these a bit more
For now, here’s just a big list of everything

## Slide 28: Functions: setup
__init__: Create everything and start up literally everything
draw_static: Helper function called in __init__ to lay out all the static elements on all the tabs. Also defines the errorloggers. Also calls load_config
stop: basically just calls stop_instruments
stop_instruments: set all valves back to the original safe position
load_config and save_config: Loads a config from or saves a new config to the local machine. When calling this any time other than upon startup, also start the ElveflowDisplay. If the default config file doesn’t exist, this crashes on startup.

## Slide 29: Functions: setup
set_oil_valve_names, set_loading_valve_names, set_cerberus_oil_valve_names, set_cerberus_loading_valve_names: load up the names
connect_to_spec: Mysterious Alex M magic (TODO: figure it out)
start_manual_thread: create the manual thread for sending commands to the microcontroller. This is a daemon thread, so we don’t need to worry about it closing
  Note that it is not a threading.thread class directly, but a custom solocomm.ManualControlThread class. Also, this thread isn’t saved as a variable? And the queue is made elsewhere?
handle_exception: redirect exceptions to the python logger

## Slide 30: Functions: exiting and graphing
save_history: I’m pretty sure this is an unused function
exit_: tell everything to shut down. Actually, this is just the elveflow_display. But there is some complicated logic with the locks to make sure of some weird race condition. I think it was the canvas.draw bug. Maybe Matplotlib and Tkinter have sorted that out?
update_graph: copy the data from ElveflowDisplay and display it on the main (Auto) tab. The logic is almost an exact copy-paste from ElveflowDisplay’s function
graph_vline: make a vertical line on the main tab’s graph

## Slide 31: Functions: buttons
auto_run_choice: start the AutoRun (first by asking the user to confirm SPEC file locations, etc.)
Several functions that put things into the command queue:
  choose_take_buffer_command, buffer_sample_buffer_command and cerberus_buffer_sample_buffer_command
  choose_take_buffer_command, take_buffer_command, and cerberus_take_buffer_command
  choose_take_sample_command, take_sample_command, and cerberus_take_sample_command
  choose_clean_and_refill_command, clean_and_refill_command, and cerberus_clean_and_refill_command
  choose_cleaning, clean_only_command, and cerberus_clean_only_command
  clean_loop
  choice_refill_only_command, refill_only_command, cerberus_refill_only_command
  load_sample_command and load_buffer_command

## Slide 32: Functions: buttons
Functions intended to be run from the command queue:
  set_refill_flag_true
  unset_purge
Several functions that put things into the manual queue:
  purge_command
  purge_soap_command
  purge_dry_command
Functions intended to be run from… somewhere
  unset_insert_purge
  unset_insert_sheath_purge
  set_insert_purge
  set_insert_sheath_purge
More functions that put things into the command queue:
  insert_purge
  insert_sheath_purge

## Slide 33: Functions: buttons
Functions intended to be run from the command queue:
  set_refill_flag_true
  unset_purge
Several functions that put things into the manual queue:
  purge_command
  purge_soap_command
  purge_dry_command
Functions intended to be run from… somewhere
  unset_insert_purge
  unset_insert_sheath_purge
  set_insert_purge
  set_insert_sheath_purge
More functions that put things into either queue:
  insert_purge (command)
  insert_sheath_purge (command)
  initialize_sheath_command (manual)

## Slide 34: Functions: UI management
toggle_sucrose
color_sucrose_button
toggle_buttons
play_done_sound

## Slide 35: Functions: hardware connection
configure_to_hardware
add_pump_set_buttons
refresh_dropdown
toggle_buttons
play_done_sound
instrument_change_values
add_pump_control_buttons
refresh_com_list

## Slide 36: Functions: hardware connection
add_rheodyne_set_buttons
AddRheodyneControlButtons
AddVICISetButtons
AddVICIControlButtons

## Slide 37: Functions: SPEC functions
ChangeDirectory
on_mkdir_timer
is_filename_safe
run_tseries: the main SPEC function. Does…something to the solocomm.controlQueue? What is that?

## Slide 38: solocomm.py
(functions outside of any class)

## Slide 39: Functions:
initConnections(MainGUI, host): make a SpecCommThread and a ControlThread (both of which are daemon threads). Returns the ControlThread.
parseListOfCommands(self, cmdList): I think this is unused. It must have gotten orphaned outside the class it’s supposed to be in, since it has a self parameter.
parse_command_file(filename): I think this is just used for testing
parseCommand(cmd): take in the cmd string and return [serv, command]
waitFor(state): every 1s, query SPEC’s state via the soloSoftCommandQueue. Once it reaches the desired state, return (blocking). The thing is, I don’t think soloSoftCommandQueue actually does anything anywhere?

## Slide 40: Queues
All of these are instances of .SpecClient.ClosableQueue, which I think is just a normal queue but implemented again…?
soloSoftCommandQueue, soloSoftAnswerQueue: Uh… I don’t see these being used anywhere (unless they’re used somewhere outside the class?)
controlQueue, ManualControlQueue: used by the control thread and the manual control thread
adxCommandQueue is used by the SpecCommThread. I don’t know about adxAnswerQueue

## Slide 41: Class: MySpecCommand
hardware/solocomm.py

## Slide 42: MySpecCommand
Extends SpecCommand.SpecCommandA
I think this is just a communication device that communicates with Spec.
The functions in here are all over the place and don’t seem to do what their names indicate. If there’s Spec documentation, probably just look at SpecCommandA.
  That said, I’ll try to put down what these functions are doing as best as I can interpret them.

## Slide 43: Functions: stuff that makes sense
__init__(self, command, host, thread):
connected(self): sets the internal self.ready to True. (This is how to change the state, not how to read it, despite the name)
disconnected(self): set the internal self.ready to False and also stop the self.thread thread.
replyArrived(self, reply): add the reply to the internal self.reply variable. (Also print out a debug message)
GetReply(self): return self.reply
ClearReply(self): clear self.reply

## Slide 44: Functions: stuff that doesn’t make sense
getReady(self): When this function is called, return a pointer to this function itself. I have no idea what the use case for this is.
beginWait(self): do nothing but print a message and then return

## Slide 45: Three classes: command threads
hardware/solocomm.py

## Slide 46: SpecCommThread
Extends thread
A specialized thread to run all the communication with Spec
It initializes a MySpecCommand from which it actually communicates with Spec.
In the main run loop, continuously check the adxCommandQueue for commands to send to Spec, and then run them. After you run each command, wait for Spec to respond.
There is no mechanism to shut down the thread, so this must be run in daemon mode
tryReconnect is never called with TryOnce=False

## Slide 47: ControlThread
Extends thread
A specialized thread to run the main communication with the microcontroller (valves/pumps) and also ADX commands.
  ADX commands are run by throwing them into the SpecCommThread queue (adxCommandQueue). It does not check to see how Spec is doing and just barrels on with the next command.
In the main run loop, continuously check the controlQueue for commands to send to Spec, and then run them. After you run each command, wait for Spec to respond.
Quits if the main Gui’s listen_run_flag() is not True; has cleanup logic when it quits

## Slide 48: ManualThread
Extends thread
Identical to ControlThread, but checks ManualControlQueue (instead of controlQueue), and does not support ADX (Spec) commands
Quits if the main Gui’s listen_run_flag() is not True; has cleanup logic when it quits

## Slide 49: FlowPath
widgets/FlowPath.py

## Slide 50: FlowPath
This really is just a graphing utility. The update functions are things like flowpath.valve2.set_auto_position("Run"), where you have to name the instance of the internal class (like valve2).
With this in mind, the internal functions should be mostly self-explanatory

## Slide 51: Whose Code Is It Anyway?
The game where everything’s made up and the points don’t matter

## Slide 52: Whose Code Is It Anyway?
gui.py: everyone
hardware/
  __init__: empty
  FileIO: Derrick
  FileIOSpoofer: Derrick—FileIOSpoofer is used for testing only
  findports: Derrick—unused
  SAXSDrivers: Josue
  solocomm: Alex M
  SPEC: Alex M—seems to be unusued, supplanted by solocomm

## Slide 53: Whose Code Is It Anyway?
widgets/
  __init__: empty except for imports
  COMPortSelector: Josue (very sparse file)
  ConsoleUi: Alex M (mostly errorlogger stuff)
  ElveflowDisplay: Derrick
  FlowPath: Alex M
  MiscLogger: Alex M (very sparse file)
  TextHandler: Derrick (very small file)
All other folders: library code from elsewhere (or images/output/other noncode)
  test folder is a failed attempt at testing